Why events localized in the center?

---
Ph2, Am241, run 537

---
Ph2, Am241, run 537

---
By Maximo
XY on s2_max_chan for Am241

---
S2 spectrum: right tail
S2 for different XY cuts:
central, edges, corners

---
Ph2, Am241, run 537

---
Ph2, Am241, run 537

---
Ph2, Am241, run 537

---
Ph2, Am241, run 537

---
S1, S2, S2/S1 for different XY cuts:
central, edges, corners

---
Ph2, Am241, run 537

---
Ph2, Am241, run 537

---
Ph2, Am241, run 537

---
Ph2, Am241, run 537

---
Mean, sigma and sigma/mean for S1 spectra for different XY cuts and runs:

---
S2/S1 for different corners

---
Ph2, Am241, run 537
(redefinition of corners)

---
Ph2, Am241, run 537
(redefinition of corners)

---
Ph2, Am241, run 537
(redefinition of corners)

---
Ph2, Am241, run 537
(redefinition of corners)

---
Luciano: did I get wrong, or there is a corner having a S2/S1 very different from the other three?

---
Am241 spectrum from F4 (bottom)

---
Marco: Here in Madrid, prompted by Davide, I made this one plot which also cast some doubts.

---
Hypotheses:
1) Luciano: Firstly, we should check to rule out "stupid" bugs in the
Reconstruction.
2) Nicola: From the corresponding laser run it seems that this channel doesn't have a different gain. So it looks like it is a like collection problem.
3) Luciano: could it be that one (or more) SiPM in the bottom tile is dead? This wouldn't affect the SER, but only the total charge spectrum
4) Nicola: George has just told me that if one SiPM dies, it can either produce a short circuit or an interruption. In both cases, he says, we must see some strong change in the gain due to the new partition of the over-voltage.

---
Hypotheses:
5) Marco: Vlad in the attachment shows that the distinct response of the "rogue" bottom tile channel is dependent on the electric field in the channel, is worse by far with no Field (run 511), and become similar to the other with increasing field (537 ~ 200 V/cm, 542 ~ 400 V/cm, 544 ~800 V/cm). Similar trend (on another scale) might be guessed for the top tile.
6) Nicola: Anyway the presence of some problematic SiPM proposed by Luciano may be a plausible explanation. When increasing the HHV, maybe some instability in the SiPM was removed.
7) Marco: it seems to me that the strange channel has a particularly long time constant, that might in principle affect indeed the integral in the fixed roi.

---
Luciano:
If the bottom effect is confirmed, I wonder: if we "bottomize" the top
tile, i.e. sum up the SiPM in groups of six, will we see a similar
effect also in the top tile? If it comes from a xy non-uniformity, I
expect to see the same trend.

---
Ph1, Am241, run 399

---
Ph1, Am241, run 418

---
Ph1, Am241, run 511

---
Ph2, Am241, run 537

---
Ph2, Am241, run 542

---
Ph2, Am241, run 544

---
By Simone
By Vlad

---
Tasks:
1) Tdrift spectrum, XY slices for Tdrift, Tdrift spectrum for XY cuts
2) XY for S2 spectrum cuts
3) Repeat all tasks using S2_max_channel algorithm